## VCSEL系統規格
規格書：型號VSO-100 Vcsel量測系統規格
積分球體(Labsphere)
尺寸：6 inch(15cm)
入射孔徑：1 inch
內部Coating：Spectralon
反射頻譜：250nm - 2500nm
反射率：200nm-250nm(R>90%)，260nm-1000nm(R>97%)
Vcsel LD峰波長及半波寬測試
光譜儀廠牌： Otsuka LE-5400
波長範圍：800-1000nm
波長解析度：0.2nm
波長精度：±0.3nm
光譜儀探測器類型：1024線陣製冷CCD
溫度變化測試波長飄移
Te-cooler廠牌：arroyo instruments 234
Temperature Range -5°C to +85°C
Recommended Controller 5305 or 585 TECSource
Pin Length Range 5mm to 9mm
峰波長變化與溫度關系 : ( 溫升1度或溫降1度 可測得 峰波長變化)
pulse width量測
廠牌：Labsphere
型號：LFPM-200-2CH
高速探測器取樣頻率範圍:Max 200kHz(可依照客戶需求更換高規)
高速探測器測量範圍:0.01~100,000mw
高速探測器探射光電流:10pA-10mA
高速探測器有效讀數:5位
高速探測器數據緩衝:192K
高速探測器線性度:0.005
高速探測器測試精度:+／-1%
高速探測器接口:USB 2.0
Power
廠牌：Keysight B2901A
最小電源解析度：1 pA /1 μV；最小量測解析度：100 fA/100 nV
最大輸出：210 V，3 A 直流/10.5 A 脈衝電流
具備任意波形產生和數位轉換功能，量測間隔為 20 µs
脈衝量測：6V/10.5A，50us至1ms

## 積分球規格
Spectralon材料特點
Spectralon是採用高温燒製的塊狀材料，採用機械加工的形式做成球形，不同於市場噴塗的工藝
耐紫外，長久照射不會發黃
漫反射率特性，接近於理想Lambertian
耐高温，可以容忍400度高温
激光損傷值高，高達8J/cm2
200nm-250nm可達90%-95%的漫反射率
無污染的情况下，該材料不會發生衰减

## 光譜儀規格書
Otsuka electronics
LE-5400
Model | standard type
Spectral range | 800nm-1000nm
Detector | Back-thinned frame transfer CCD
Number of pixels | 1024×58
Spectral resolution *1 @546.07nm | 0.2nm
Data point interval | 0.02nm
Wavelength accuracy | ±0.3nm
Stray light (broadband for illuminant A) *2 | 5・10E-4
Stray light for LEDs *3 | 1・10E-4
Integration time | 2msec-20sec
Linearity | ±0.5%
Cooling | +5℃
Spectroradiometry
Irradiance sensitivity range *4 | TBD
Signal sensitivity at 1s integration time *4 | -
spectroradiometric accuracy *5 | ±3.5%
Spectrophotometry
Bseline noise *6 | ±0.4%
Transmission measuring accracy | -
Baseline drift *7 | 0.2%/h
Measurement uncertainty (for LED measurement)
Dominant wavelength | ±0.5nm
Color coordinates (x,y) *9 | ±0.002
Spectrograph
Focal length, grating | Approx. 135mm f/3 concave grating
Slit | Standard：50μｍ
Filter wheel | not available
Electrical data
AD converter | 16-bit resolution
PC interface | USB2.0
Triggering | input : Isolated input (12V to 24V) Start,Dark1,Dark2 output : Isolated ioutput(12V to 24V)BUSY,INMEAS,SYNCRO,OVER,UNDER
Other
Dimensions(H,W,D) [mm] | 160×280×296
Power consumption | ｍax.100VA
Ambient temperature | 20-35℃；0-80% non condensing
Weight | approx.9.2ｋｇ
Applicable standards | Meets EN61326:2006、EN61010-1:2001(2nd Edition)

## Te-cooler規格
234 TEC TO-Can LaserMount |  | Description | 585-04-08 | 585-05-12
Laser |  | Drive Channel
Packages Supported | 5.6mm and 9mm TO-Can, 3 & 4 pin | Current
| Flange mount, 12mm or 18mm | Range (A) | 4 | 5
Fiber Pigtailed Package Support | horizontal hole spacing, or 0.75” | Compliance Voltage (V) | 8 | 12
| (19.1mm) vertical hole spacing. | Max Power (W) | 28 | 55
Pin Length Range | 5mm to 9mm | Resolution (A) | 0.01
Maximum Laser Current | 3 Amps | Accuracy (+/- [% set point + mA]) | 0+30
Temperature Control |  | Noise/Ripple (mA, rms) | <5
Temperature Range | -5°C to +85°C | Temperature Control
Sensor Type | 10kΩ Thermistor | Range (ºC)1 | -99 to 250
TE Module (at 25°C) | Imax = 5.6A | Resolution (ºC) | 0.01
| Vmax =7.2V | Thermistor Accuracy (± ºC)2 | 0.05^3
Recommended Controller | 5305 or 5240 TECSource | Short Term Stability (1hr) (± ºC)4 | 0.004
Mechanical Summary 234 |  | Short Term Stability (24hr) (± ºC)4 | 0.01
Center hole | 1.035”-40, threaded through | Measurement Channels
Cage mount | 4-40 x 4 holes, 30mm on center | Current
Size (H x W x D) [in(mm)] | 3.0 (76.2) x 3.0 (76.2) x 2.91 (73.9) | Resolution (mA) | 10
234 with 234-C-02 Cover Plate |  | Accuracy (+/- [% reading + mA]) | 0+30
Center hole 0.27” | [6.9mm], straight through | Voltage
Cage mount | 4-40 x 4 holes, 30mm on center | Resolution (mV) | 10
Size (H x W x D) [in(mm)] | 3.0 (76.2) x 3.0 (76.2) x 3.19 (80.9) | Accuracy (+/- [% reading + mV]) | 0+50
Connectors |  | Sensor
Laser Diode | DB-9, male | 10μA Thermistor
Mount TEC | DB-15, male | Range (kΩ) | 0.1 - 450
Nitrogen | 1/16” barb | Resolution (Ω) | 10
Case Ground | 4mm banana jack | Accuracy (+/- [% reading + Ω]) | 0.05 + 50
General |  | 100μA Thermistor
Fan Power | 12VDC, 210mA max | Range (kΩ) | 0.05 - 45
Mounting holes | 8-32 and M4 threaded holes | Resolution (Ω) | 1
|  | Accuracy (+/- [% reading + Ω]) | 0.05 + 5